## Method Information

Restek Stabliwax is better than HP INNOWax in FAME analysis.

Columns with 0.32 µm D.I. is better than 0.23 or 0.53 µm.

## Method Change History

Operator	Date	Change Information

Wei Zou	5/7/01

## Run Time Checklist

Pre-Run Cmd/Macro:	off

Data Acquisition:	on

Standard Data Analysis:	on

Customized Data Analysis:	off

Save GLP Data:	on

Post-Run Cmd/Macro:	off

Save Method with Data:	off

## Injection Source and Location

Injection Source:	Manual

Injection Location:	Front

===============================================================

#### HP6890 GC METHOD

===============================================================

## OVEN

Initial temp:	200 'C (On)	Maximum temp:	250 'C

Initial time:	1.00 min	Equilibration time:	1.00 min

Ramps:

#	Rate		Final temp	Final time

1	10.00		210		0.10

2	5.00		220		0.10

3	10.00		240		8.00

4	0.0(Off)

Post temp:		200 'C

Post time:		0.50 min

Run time:		14.20 min

FRONT INLET (SPLIT/SPLITLESS)	BACK INLET (SPLIT/SPLITLESS)

Mode:		Split				Mode:	Split

Initial temp:	250 'C (On)			Initial temp:	250 'C (On)

Pressure:	24.4 psi (On)			Pressure:	3.3 psi (On)

Split ratio:	35:1				Split ratio:	35:1

Split flow:	48.5 mL/min			Split flow:	157.5 mL/min

Total flow:	52.2 mL/min			Total flow:	164.5 mL/min

Gas saver:	Off				Gas saver:	Off

Gas type:	Helium				Gas type:	Helium

COLUMN 1	COLUMN 2

Capillary Column	Capillary Column

Model Number:	HP 19091N-133	Model Number:	Restek Stabliwax

HP-INNOWax Polyethylene Glycol

Max temperature:	260 'C	Max temperature:	250 'C

Nominal length:	30.0 m	Nominal length:	15.0 m

Nominal diameter:	250.00 um	Nominal diameter:	530.00 um

Nominal film thickness:	0.25 um	Nominal film thickness:	1.00 um

Mode:	constant pressure	Mode:	constant pressure

Pressure:	24.4 psi	Pressure:	3.3 psi

Nominal initial flow:	1.4 mL/min	Nominal initial flow:	4.5 mL/min

Average velocity:	38 cm/sec	Average velocity:	49 cm/sec

Inlet:	Front Inlet	Inlet:	Back Inlet

Outlet:	Front Detector	Outlet:	Back Detector

Outlet pressure:	ambient	Outlet pressure:	ambient

FRONT DETECTOR (FID)	BACK DETECTOR (FID)

Temperature:	275 'C (On)	Temperature:	275 'C (On)

Hydrogen flow:	40.0 mL/min(On)	Hydrogen flow:	40.0 mL/min(On)

Air flow:	450 mL/min(On)	Air flow:	450 mL/min(On)

Mode:	Const. col.+makeup	Mode:	Const. col.+makeup

Combined flow:	50.0 mL/min (On)	Combined flow:	50.0 mL/min (On)

Makeup flow:	30.0 mL/min (Off)	Makeup flow:	30.0 mL/min (Off)

Makeup flow:	On	Makeup flow:	On

Makeup Gas Type:	Helium	Makeup Gas Type:	Helium

Flame:	On	Flame:	On

Electrometer:	On	Electrometer:	On

Lit offset:	2.0	Lit offset:	2.0

SIGNAL 1	SIGNAL 2

Data rate:	50 Hz	Data rate:	50 Hz

Type:	front det - col comp 1	Type:	back det - col comp 2

Save Data:	On	Save Data:	On

Zero:	0.0 (Off)	Zero:	0.0 (Off)

Range:	0	Range:	0

Fast Peaks:	Off	Fast Peaks:	Off

Attenuation:	0	Attenuation:	0

## COLUMN COMP 1	COLUMN COMP 2

Derive from front detector	Derive from back detector

## POST RUN

Post Time:	0.50 min

Oven Temperature:	200 'C

Column 1 Pressure:	24.4 psi

Column 2 Pressure:	3.3 psi

## TIME TABLE

Time	Specifier	Parameter & Setpoint

0.00	Signal 1:	front det - col comp 1

0.00	Signal 2:	back det - col comp 2

===============================================================

#### Calibration Table

===============================================================

WEI_FA

Calib. Data Modified:	1/15/01 3:57:56 PM

Calculate:	External Standard

Based on:	Peak Area

Rel. Reference Window:	4.000 %

Abs. Reference Window:	0.000 min

Rel. Non-ref. Window:	4.000 %

Abs. Non-ref. Window:	0.000 min

Uncalibrated Peaks:		not reported

Partial Calibration:		Yes, identified peaks are recalibrated

Correct All Ret. Times:	No, only for identified peaks

Curve Type:	Linear

Origin:	Forced

Weight:	Equal

Recalibration Settings:

Average Response:		Average all calibrations

Average Retention Time:	Floating Average New 75%

Calibration Report Options:

Printout of recalibrations within a sequence:

Calibration Table after Recalibration

Normal Report after Recalibration

If the sequence is done with bracketing:

Results of first cycle (ending previous bracket)

Signal 1: FID1 A,

RetTime(m)	Lvl	Amt[ng]	Area	Amt/Area	Grp Name

3.270	1	3000	1331	2.25297	C16:0

2	6000	2406	2.49354

3	9000	3611	2.49221

3.900	1	3000	1275	2.35289	C17:0

2	6000	2440	2.45820

3	1.50000e4	5249	2.85720

4.600	1	3000	1321	2.26963	C18:0

2	6000	2387	2.51318

3	9000	3590	2.50682

4.870	1	3000	1371	2.18685	C18:1

2	6000	2482	2.41661

3	9000	3718	2.42012

5.200	1	3000	1353	2.21676	C18:2

2	6000	2448	2.45097

3	9000	3677	2.44751

5.370	1	30	11.85	2.53079	C19:0

2	75	28.66	2.61625

3	150	60.05	2.49792

5.800	1	3000	1323	2.26728	C18:3

2	6000	2395	2.50439

3	9000	3601	2.49903

7.440	1	30	12.01	2.49750	C20:4

2	60	23.58	2.54367

3	100	41.01	2.43825

8.300	1	50	16.73	2.98846	C20:5

2	100	33.38	2.99509

3	200	71.71	2.78878

11.400	1	50	14.98	3.33689	C22:5

2	100	31.18	3.20626

3	200	69.56	2.87513

12.000	1	30	9.25	3.24184	C22:6

2	60	19.14	3.13349

3	100	35.28	2.83423